## System Requirements
Requirement Type | Number | Requirement | Verification Method
Design | 1.1 | The system shall accomodate a 12x12" PCB as device to be tested. | Demonstration
| 1.2 | The system shall include insulation designed for external safety according to ASTM C1055. | Inspection
| 1.3 | The system's design shall be safe for continual operation according to NFPA-86. | Inspection
System | 2.1 | The system shall include a port to allow electrical cables to enter and leave for stimulus and measurement of the PCB. | Inspection
| 2.2 | The system shall comply with electrical grounding requirements. | Inspection
| 2.3 | The system shall be able to operate on a standard wall-outlet power of 120V (AC). | Demonstration
| 2.4 | The system will comply with appropriate NEC and UL rules governing powered and plug in devices. | Inspection
Functional | 3.1 | The system shall monitor temperature inside the chamber. | Demonstration
| 3.2 | The system shall control temperature inside the chamber. | Demonstration
| 3.3 | The system shall remain within a specified humidity range (0% - 90%). | Demonstration
| 3.4 | The system shall provide data output. | Demonstration
Interface | 4.1 | The system shall include an HMI for the control system. | Inspection
| 4.2 | The system shall include separate operating/control modes. | Demonstration

## Subsystem Requirements
Subsystem | Number | Requirement | Performance | Margin | Verification Method | Basis / Notes | Other Possible Sources:
Mechanical (Structural) | STR-01 | The system shall have external base dimensions of no more than 30x30 inches. | 30x30 in | -2x2 in | Inspection | Analysis | https://store.astm.org/products-services/standards-and-publications/standards/electrical-insulating-material-standards.html
| STR-02 | The system shall have internal base dimensions of no less than 12x12 inches. | 12x12 in | +2x2 in | Inspection | Analysis
| STR-03 | The chamber opening(s) shall include proper thermal sealing according to ASTM F1886. | Complies | N/A | Inspection | Maybe: ASTM D5470-17 Test procedure for thermal calculations of seal
| STR-04 | The internal thermal control components shall include proper thermal sealing according to ASTM C633-01. | Complies | N/A | Inspection | Maybe: ASTM D5470-17 Test procedure for thermal calculations of seal
Mechanical (Thermo-fluid) | THE-01 | The system shall cover commercial temperature range (0oC - 70oC) | 0-70oC | +15oC, -40oC | Measure | Analysis
| THE-02 | The system shall be capable of sweeping all temperature ranges within a 24-hr cycle. | <24hr | -1hr | Test | Analysis
| THE-03 | The heating system should be contained/unexposed within the system. | Complies | N/A | Inspection
| THE-04 | The cooling system should be contained/unexposed within the system. | Complies | N/A | Inspection
Electrical (Component) | COM-01 | The PCB shall be operatable while temperature is controlled. | Complies | N/A | Measure
| COM-02 | System will adhere to NEC rules on grounding according to Article 250. | Complies | N/A | Inspection | Regulations
| COM-03 | System will adhere to NEC rules on appropriate fuses and load wire sizes. | Complies | N/A | Inspection | Regulations
| COM-04 | System will adhere to UL 61010 for test and measurement equipment. | Complies | N/A | Inspection | Regulations
| COM-05 | The sensors shall be properly rated to withstand internal environmental conditions. | Complies | N/A | Analysis
| COM-06 | The system shall support PC connection for datalogging. | Complies | N/A | Inspection
Electrical (Control System) | CON-01 | The system shall allow local temperature setpoint control. | Complies | N/A | Test
| CON-02 | The system shall allow remote temperature setpoint control. | Complies | N/A | Test
| CON-03 | The system shall maintain a toleranced temperature at ±3oC of setpoint. | ±3oC | ±0oC | Measure | Analysis
| CON-04 | The system shall prohibit thermal control when the chamber door is open. | Complies | N/A | Demonstration
| CON-05 | The system shall prohibit thermal control when a user is handling the PCB. | Complies | N/A | Demonstration
| CON-06 | The system shall log temperature data. | Complies | N/A | Test
| CON-07 | The system shall log humidity data. | Complies | N/A | Test
| CON-08 | The system shall log faults for safety measures. | Complies | N/A | Measure
| CON-09 | The system shall have a "shutdown" state. | Complies | N/A | Demonstration
| CON-10 | The system shall have a "standby" mode. | Complies | N/A | Demonstration
| CON-11 | The system shall have a normal operating "start" mode. | Complies | N/A | Demonstration
| CON-12 | The system shall allow for manual switching between modes. | Complies | N/A | Demonstration
Safety | SAF-01 | The internal surfaces shall be safe for electronics according to ASTM D3874-20. | Complies | N/A | Analysis
| SAF-02 | The system external surfaces shall not exceed the safe-touch temperature of 50oC according to ASTM C1055. | 50oC | -7oC | Analysis | Analysis
| SAF-03 | The system shall be able to operate in normal environmental (external) conditions of 21oC. | 21oC | ±5oC | Demonstration | Analysis
| SAF-04 | The system shall include a safety measure if temperature range & tolerance is exceeded. | Complies | N/A | Demonstration
| SAF-05 | The system shall have a "safe" state. | Complies | N/A | Demonstration
Note on C4: :: ASTM F1886/F1886M-16 outlines visual inspection test procedure for seal integrity.
Note on C28: :: ASTM D3874-20 outlines standard test procedure for ensuring safety of electronic components (low risk of ignition).

## Other Tables
Requirement Types | Examples
Design | Structural, Stiffness, Materials, Reliability, Redundancy, Electronic Parts
Environmental | Temperature, RH, Pressure, Vibration, EMC/EMI, Cleanliness
Functional | Operating Modes, Software, Mechanism Control, Autonomous Mode
Interface | Power Supply, Data Handling, Mounting, Thermal
Performance | Resolution, Focus, Sampling Rate, Jitter Control
Project | Program Management, Systems Engineering, Quality Assurance
Resource | Mass, Weight, Physical Size, Power, Data Volume, Data Rate, Field-of-View
Safety | Personal Protective Equipment, Machine Guards
Verification | Models, Methods, Levels
Verification Method | Description
Analysis | By mathematically modeling or simulating the system, it is predicted that a requirement is met or not.
Demonstration | By doing something with the system, it is clear if the requirement is met or not.
Inspection | By observing the system, it is clear if the requirement is met or not.
Measure | By collecting multiple measurements, it is clear if the requirement is met or not.
Test | By collecting a single measurement, it is clear if the requirement is met or not.
Subsystem | Classification
Mechanical (structural) | STR
Mechanical (thermo-fluid) | THE
Electrical (component) | COM
Electrical (control system) | CON
Safety | SAF
Basis | Description
Analysis
Research
Regulations
Customer

## Project Risks
Risk Matrix | [9 empty] | Consequence Category | PRE-MITIGATION | [6 empty] | POST-MITIGATION
Consequence | 5 | B, L | A, E, P | K, M, O | Z |  | Letter | Consequence | Description | Asset
| 4 |  | Y | N | I | Q, W | A | Asset | Mechanical Failure of Components (worn, broken) | Cost | Consequence | 5 |  |  |  | Z |  | Consequence | 5
| 3 |  | F, G | T, U, X | V |  | B | Asset | Insufficient Dimensions for PCB testing (internal, external, cable port, chamber door) | Human Safety |  | 4 |  |  |  |  | Q, W |  | 4 |  | Z
| 2 | R | J |  | C, D | H | C | Asset, Human Safety | Internal System is Over Temperature Range (fire, melting components) | Performance |  | 3 | [6 empty] | 3 |  |  |  | Q
| 1 | S |  |  |  |  | D | Asset, Human Safety | Internal System is Under Temperature Range (freezed components) | Schedule |  | 2 | [6 empty] | 2 |  |  |  | W
|  | 1 | 2 | 3 | 4 | 5 | E | Asset, Performance | System's Safety Measure Fails If Outside Temperature Range |  |  | 1 | [6 empty] | 1
|  | Likelihood |  |  |  |  | F | Asset, Performance | Internal Heating/Cooling System Can't Reach upper/lower boundaries (poor components, inefficient controller) |  |  |  | 1 | 2 | 3 | 4 | 5 |  |  | 1 | 2 | 3 | 4 | 5
Notes: | Combined Risk Scores: 0-9 (Green) --> Minor 10-19 (Yellow) --> Moderate 20-29 (Red) --> Major | [5 empty] | G | Asset, Performance | Inconsistent Temperature Control (poorly designed microcontroller or interface) |  |  |  | Likelihood | [6 empty] | Likelihood
[7 empty] | H | Asset, Performance | Failure to Maintain Toleranced Setpoint (±o3C) (ineffective control system, poor insulation - internal & external) |  | Notes: | Combined Risk Scores: 0-9 (Green) --> Minor 10-19 (Yellow) --> Moderate 20-29 (Red) --> Major | [5 empty] | Notes: | Combined Risk Scores: 0-9 (Green) --> Minor 10-19 (Yellow) --> Moderate 20-29 (Red) --> Major
[7 empty] | I | Asset, Performance | Lack of Electronic Protection (excessive moisture buildup)
[7 empty] | J | Cost | Spending Past Designated Budget for Equipment & Materials
[7 empty] | K | Human Safety | Lack of Human Protection at Critical Temperature Ranges (burns)
[7 empty] | L | Human Safety | System Fails to Meet Environmental/Engineering Codes
[7 empty] | M | Human Safety | System Fails to Prohibit Temperature Control During Manual User Usage.
[7 empty] | N | Human Safety, Performance | Chemical Leaking
[7 empty] | O | Human Safety, Performance | Electrical Hazards (short circuits, improper connections)
[7 empty] | P | Human Safety, Performance | System Fails to Operate Safely Under Normal Conditions
[7 empty] | Q | Performance | Software/HMI Failure
[7 empty] | R | Performance | Lack of Refrigerant Fluid in System
[7 empty] | S | Performance | Controller not User Friendly (poor HMI design)
[7 empty] | T | Performance | Sensor Reading Failure (miscalibration, component issue)
[7 empty] | U | Performance | Improper Ventilation (pressure changes, fumes buildup)
[7 empty] | V | Performance | Ineffective Insulation (poor materials or design)
[7 empty] | W | Performance | Data Measuring, Collecting, and Logging Fail to Work Systematically
[7 empty] | X | Performance | System Mode(s) Failures
[7 empty] | Y | Schedule | Poor Scheduling for Fabricated Parts or Sourced Materials
[7 empty] | Z | Schedule, Cost | Rework of Product (ineffective design from failed systems/components/materials)

## Verification Matrix (Example)
Requirement Type→ | Performance |  |  | Functional |  |  | Design |  | Interface | Resource |  | Verification Reference
Requirement Title→ | Measure environmental & compost conditions | Connects to system | Meets consistent system power requirements | Indicate threshold reached | Report data to Website | Hold up to 200 kg | Environmentally Friendly | Reliable & Durable | User Friendly | Meets the required budget | Efficient use of materials and resources | This column identifies the specific verification summary related to the activity on each line. Multiple summaries may apply to a single line item.
Level of Assembly
Systems
Arduino (electronics assembly) |  | D | M |  | T |  |  |  | A | V
Sensor Container (housing assembly) | [6 empty] | A | A | A | V | A
Door Mechanism |  | D |  |  |  |  | A | T | T | V | A
Components
MicroController | M |  |  | T |  |  |  |  | D | V
HMI | M |  |  | T |  |  |  |  | D | V
| M |  |  | T |  |  |  |  | D | V
Materials
Refrigerant | M |  |  | T |  |  |  |  | D | V
Metal Housing | M |  |  | T |  |  |  |  | D | V
| M |  |  | T |  |  |  |  | D | V
Verification Method: | Notes:
A - Analysis | 1. YELLOW indicates tasks to be done. GREEN indicates completed tasks.
D - Demonstration
I - Inspection
M - Measurement
T - Test